www.china-suke.com

1. SUMMARY

1.1Features

· Measurement is not affected by the variation of flow density, viscosity, temperature, pressure and conductivity. High accuracy measurement is guaranteed according to the linear measurement principle.

· No obstacle in the pipe, no pressure-loss and lower requirement for straight pipeline.

· DN 6 to DN2000 covers a wide range of pipe size. A variety of liners and electrodes are available to satisfy different flow characteristic.

· Programmable low frequency square wave field excitation, improving measurement stability and reducing power consumption.
· Implementing 16 bits MCU, providing high integration and accuracy; Full-digital processing, high noise resistance and reliable measurement; Flow measurement range up to 1500:1.

· High definition LCD display with backlight.

· RS485 or RS232 interface supports digital communication.

· Intelligent empty pipe detection and electrodes resistance measurement diagnosing empty pipe and electrodes contamination accurately.

· SMD component and surface mount technology (SMT) are implemented to improve the reliability.

1.2Main Applications

  electromagnetic flowmeter can be used to measure the volume flow of conductive fluid in a closed pipeline. It is widely applied in the flow measurement and control in the fields of chemical and petroleum industry, metallurgy industry, water and waste water, agriculture and irrigation, paper making, food and beverage industry and pharmaceutical industry.

1.3Ambient Conditions

Ambient temperature:  sensor: -25℃ to + 60℃; converter: -25℃ to + 60℃.

Relative humidity: 
  5% to 90%;

1.4Working Conditions

Maximum fluid temperature: 


Compact type: 60℃
Remote type: 
Teflon   
150℃
Neoprene

80℃; 120℃
Polyurethane
70℃

 Fluid conductivity: ≥ 5S/cm

2. WORKING PRINCIPLES
2.1 Measuring Principles

The measuring principle of electromagnetic flowmeter is based on the electromagnetic induction law of Farady. The sensor is mainly composed of measuring tube with isolate lining, a pair of electrodes installed by penetration of the measuring tube wall, a pair of coils and iron core to produce working magnetic field. When the conductive fluid flows through the measuring tube of the sensor, the voltage signal in direct proportion to the average flow velocity of the fluid will be inducted on the electrodes. The signal is amplified and treated by the transmitter to realize various display functions. 
                 [image: image1.jpg]Converters

Figure one:Principle of Electronagnetic Flometer




 2.2 Converter Circuit Schematic
The converters supplies a stable exciting current to the coil in the sensor of electronetic flowmeters to get B constant and amplifies the electromotive force and convert it into standard signals of current or frequency so that the signals can be used for displaying, controlling and processing. The schematic of converter circuit is shown in Fig. 2.1. 

[image: image2.png]]

1-50008z

Frequency or
Pulse Output

0¢ Gate
Status output

Rsies

erting

circuit

16 bit
o

STTching
PoserSuppl

Trrent
output.

Pulse Output

oo
Contral

[CommmTcatTor]

Interface

won

EEROn

™
Displa

Keyboard





Fig. 1 Schematic of converter circuit

3. PRODUCT CLASSIFICATION
3.1Product components

Electromagnetic flowmeter is composed of sensor and converter. The remote-type also needs a special double-layer shielded cable to connect the converter and the sensor.

3.2Product models

 electromagnetic flowmeter has two forms: compact-type and remote-type. There are seven kinds of electrode materials and four kinds of lining materials available for sensor.

4. TECHNICAL SPECIFICATION 
4.1The flowmeter is compliant to Standard “JB/T 9248-1999 Electromagnetic Flow Meter”.
4.2Maximum flow speed: 15m/s

4.3DN size: 3, 6, 10, 15, 20, 25, 32, 40, 50, 65, 80, 100, 150, 200, 250, 300, 350, 400, 450, 500, 600, 700, 800, 900, 1000, 1200, 1400, 1600, 1800, 2000
4.4Accuracy: 0.5%, 0.2%

4.5Nominal Pressure:
4.0MPa(DN0-150)

1.6MPa(DN200-600)

1.0MPa(DN700-1200)

0.6MPa(DN400-2000)

Or other specified by order

4.6Materials

Electrode Forms and Materials:

Electrode design has four forms: standard form, scraper form, detachable form and grounding electrode.

Electrode material is selectable from stainless steel containing Mo, stainless steel coated with carbonized tungsten, Hastelloy B, Hastelloy C, Titanium, Tantalum and platinum-iridium alloy.

Flange material: carbon steel

Grounding ring: stainless steel

Inlet protection ring: carbon steel, stainless steel

4.7Enclosure

IP65:

IP68: only available for remote-type sensor with neoprene or polyurethane liner, not including flameproof model.

4.8Flameproof Standard

DN15-DN600 compact-type: md II BT4

DN15-DN1600 remote-type: sensor and converter both installed in dangerous area: md II BT4.

DN15-DN1600 remote-type: sensor installed in dangerous area and converter installed in safe area: md II BT4

4.9Connection Cable

Special cable is needed to connect the sensor and the converter for remote-type flowmeter. The cable length should not exceed 100 meters. A 10-meter cable is supplied free of charge and the rest is to be ordered.

4.10General Specification of Converter 

· Power supply:
AC 85-265V, 45-63Hz, ≤20W;  DC 11-40V

· Converter display and operation: four keys are available to set all parameters. External handheld or PC can be used to do the configuration. High resolution LCD display with backlight, empty pipe detection and self-diagnostic function are equipped in the converter.

· Digital communication: RS485, RS232, MODBUS,REMOTE

· Output Signals: 

- Current output: 
fully-isolated, 4-20mA/0-10mA

load resistance: 0-10mA: 0-1.5KΩ; 4-20mA: 0-750Ω.


- Frequency output: bidirectional flow output. Frequency output is proportional to the flow percentage of the full range. The converter provides fully isolated transistor open collector frequency output ranged from 1 to 5000 Hz. The external DC power supply should not exceed 35V and maximum collector current is 250mA. 

- Pulse output: bidirectional flow output. The converter can output up to 5000cp/s pulse series, which is dedicated to external totalization. Pulse factor is defined as volume or mass per pulse. It can be set to 0.001L/p, 0.01L/p, 0.1L/p, 1L/p, 2L/p, 5L/p, 10L/p, 100L/p, 1m3/p, 10 m3/p, 100 m3/p or 1000 m3/p. Pulse width is selectable from auto, 10ms, 20ms, 50ms, 100ms, 150ms, 200ms, 250ms, 300ms, 350ms and 400ms. Photo-coupler isolated transistor open collector circuit is used for pulse output. The external DC power supply should not exceed 35V and maximum collector current is 250mA.

- Flow direction indication: The converter is capable of measuring both forward and reverse flow and recognizing its direction. The converter outputs 0V low level for forward flow, while +12V high level for reverse flow.

- Alarm output: Two channels of photo-coupler isolated open collector circuit are used for alarm signal output. There are two alarm outputs: high limit alarm and low limit alarm. The external DC power supply should not exceed 35V and maximum collector current is 250mA.
- Damping constant: Damping time is selectable from 0.2 to 100s.

5. APPEARANCE AND INSTALLATION
Dimensions of Converter, shown as Fig 3.

[image: image18.jpg]214

164

3-086.5

-066.08
40000000415
]|

BELS:4T0 My

234

E=(b) HHBEBINE

123

e

2415

E=(b) EFRBINIFT





[image: image19.jpg]Appropriate Overhang

A
W

Strong Sunshine Prevention

Keeping Away from Leaking Keeping Far Away from Fire

The correct installation flowmeter



[image: image20.jpg]Electrode axial keeps horizonta

Em

Support
t——— ——

Reasonable Support

I 5xDN | 3xDN

Ensure. the Requir. of the Strafght Pipe section

Full of Pipe

Measurement for the Precipitable

Easy to Maintenance and Clean—up

Not Installed in Front of the Inlet of Pump

Strong Shake Avoidance





[image: image21.jpg]Connerting Wine

2\

[

:\\

Grourid Resistance<10Q

Earth ==

<400

Metal Pipes

>4mm? Copper

Sealing gaske

Sea}ing

>4mm2Copper

Grourid Resistance <10Q
Barth= <40Q

Non-metallic pipes, grounding electrode sensors installed

%

Link
5 1 Grounding Flange
gasket‘s
1 h Flow
——

: >4mm2Copper
Earth -round Resistance<10Q

o3 <40Q

Non-metallic pipes, no grounding electrode sensors

Connerting Wine

Gr\n\mding Flafige \ >4mm? Copper
Earth~‘_—6mund Resistance<10Q
3| <40Q

Grounding Flange

With the pipeline cathodic protection





[image: image22.emf] 

AL AH P+COMI+ COM L1 L2

+

+ + + + + + + + +

+

+ + + + + + + + + + +

+

T- T+ FUSE

FDIR

+NI

-NI


[image: image23.emf] 

 

ALM+  

 

 

 

 

ALM -  

-  

 

 

ALCOM  

 

DC Power  

-  

+  

Low Alarm  

High Alarm  

[image: image3.jpg]Dimensions of Sensor, shown as Fig 4.

Remote—type Nominall ) Moprocinate Weight | |Flange Size ' (GBT9119  4.0MPa)
L w H T =
Diameter lomeet-tye{ Sensor

0 | 200 | %0 | 176 ] 95 | 65
5 |2 | o |76 | 0| 7
2 | 200 | 105 |16 | 12 | 9
25 |20 |10 | 76 | 1| 1
[0 | 200 | 150 | 210 | 16 | 18
50 | 200 | ves | 210 | 17 | 1
o5 |20 |15 [0 | %5 | 2
50 | 200 | 200 | 250 | 29 | 25

10 | 250 | 235 212 | 31 | 28
DN10-~DN150 150 | 300 m‘J_ 3 | 4t | 38

Terminal Block

Remote-type

Noninal |

Diameter]
200
250
300

,_
+

DN200~DN800

H o | koot

Weight
K

836 | 435
936 | 545

900 | 900 | 1036 | 655
1000 | 1000 | 1136 | 810
1200 | 1200| 1336 | 875
1400 | 1400 | 1536 | 1235
1600 | 1600 | 1736 | 1555

1800 | 1800 | 1560 | 2085

2000 | 2000 2160 | 2610

Remote-type

Specified Cablé”

- \L:/'Terminal Block

DN700~DN2000





6. CONVERTER OPERATION MENU AND PARAMETER SETTING

6.1 Keypad and Display
[image: image4.emf] 

282.74

 m3/h

！

∑+00123456.78 m3

Volume Flow

Flow Speed

Flow Percentage

Electrode Resistance

Forward Total Flow

Reverse Total Flow

Net Total Flow

Alarm Display

ENTER, Combined with  ALT  to enter menu

UP,Add 1, PageUp, combined with ALT to shift right

DOWN,Minus 1, PageDown, Shift left with ALT

ALT, combined with other key

Mass Flow

Alarm

Symbol

Flow Range Ratio


Fig.5(a) Remote-type key and display

 SHAPE \* MERGEFORMAT 



Fig.5(b) Compact-type key and display

6.2 Keypad Function

1. Automatic Measurement Mode

DOWN:


 
Scroll bottom line display;

UP:




Scroll top line display;

ALT + ENTER:

Enter into setting mode;

ENTER:



Return to measurement mode.

2. Parameter Setting Mode

DOWN:



Subtract one form the digit at the cursor;

UP:




Add one on the digit at the cursor

ALT + DOWN:

Cursor shifts left

ALT + UP:


Cursor shifts right

ENTER:



Enter/exit submenu;

ENTER:



Return to measurement mode if held for 2 seconds at 





any location

Remarks: 

(1) When using ALT key, hold ALT first and then press UP or DOWN.

(2) Under setting mode, the meter returns to measurement mode automatically if no key is pressed for 3 minutes.

(3) When adjusting flow zero, UP or DOWN key can be used to change the sign (+/-).

(4) When setting flow range, UP or DOWN key can be used to change flow unit.

6.3 Parameter Configuration and Operation Password
The setting parameters determine the operation status, calculation method and output mode of the flow meter. Properly setting meter parameter can make the meter work in best condition and higher accuracy of display and output can be obtained.

There are five levels of password, where level 0 - 3 are open for user and level 4 reserved for manufacturer. Level 1 to 2 passwords are changeable by higher level password-holder, e.g. Level-3 password.

Meter setting can be browsed by entering any level of password. However, higher level password is needed to change settings.

· Password Level-0 (default value 0521): fixed and browsing only;

· Password Level-1 (default value 7206): changeable and authorized to modify menu item 1 to 25;

· Password Level-2 (default value 3110): changeable and authorized to modify menu item 1 to 29;

· Password Level-3 (default value 2901): fixed and authorized to modify menu item 1 to 38;

· Password Level-4 (reserved):
 fixed and authorized to modify any menu item including resetting system.

· Totalizer Reset Password (default value 36666): changeable in menu item ‘Clr Tot. Key ’and authorized to clear the three internal counter.

It is suggested that Level-3 password be held by manager or supervisor while Level-0 to 2 passwords be kept by operator. The Level-3 password can also be used to change the password for totalizer resetting.

6.4 Table of Parameter Setting Menu

The converter setting menu consists of 45 items. Many of them are set up by manufacturer before shipping. It is not necessary to change them when applying. There are only a few of them to be set by user according to the application.  The menu items are listed in the table below:

	Item No.
	Menu Display
	Setting Method
	Password Level
	Value Range

	1
	  Language
	Option
	1
	English

	2
	  Sensor Size
	Option
	1
	3 - 3000mm

	3
	  Flow Range
	Modify
	1
	0 - 99999

	4
	  Auto Rng Chg
	Option
	1
	ON / OFF

	5
	  Damping
	Option
	1
	0 - 100 s

	6
	  Flow Dir.
	Option
	1
	Fwd/ Res

	7
	  Flow Zero
	Modify
	1
	+/-0.000

	8
	  L.F. Cutoff
	Modify
	1
	0 - 99%

	9
	  Cutoff Enble
	Option
	1
	ON / OFF

	10
	  Rate-Of-Chng
	Modify
	1
	0 - 30%

	11
	  Limit Time
	Modify
	1
	0 - 20 s

	12
	  Total Unit
	Option
	1
	0.0001L - 1 m3

	13
	  Flow Density
	Modify
	1
	0.0000 - 3.9999

	14
	  Current Type
	Option
	1
	4-20mA/0-10mA

	15
	  Pulse Output
	Option
	1
	Frq/ Pulse

	16
	  Pulse Factor
	Option
	1
	0.001L - 1 m3

	17
	  Freq Max
	Modify
	1
	1 - 5999 Hz

	18
	  Comm Address
	Modify
	1
	0 - 99

	19
	  Baudrate
	Option
	1
	600 - 14400

	20
	  EmpPipe Det.
	Option
	1
	ON / OFF

	21
	  EmpPipe Alm
	Modify
	1
	200.0 KΩ

	22
	  Hi ALM Enble
	Option
	1
	ON / OFF

	23
	  Hi Alm Limit
	Modify
	1
	000.0 - 199.9%

	24
	  Lo Alm Enble
	Option
	1
	ON / OFF

	25
	  Lo Alm Limit
	Modify
	1
	000.0 - 199.9%

	26
	  RevMeas.Enbl
	Option
	1
	ON/OFF

	27
	  Sensor S/N
	Modify
	2 
	000000000000-999999999999

	28
	  Sensor Fact.
	Modify
	2
	0.0000 - 3.9999

	29
	  Field Mode
	Option
	2
	Mode 1,2,3

	30
	  Multiplying
	Modify
	2
	0.0000 - 3.9999

	31
	  F. Total Set
	Modify
	3
	0000000000 - 9999999999

	32
	  R.Total Set
	Modify
	3
	0000000000 - 9999999999

	33
	  Input Contrl
	Option
	3
	Disable/Stop Tot/Reset Tot

	34
	  Clr Totalizr
	Password
	3
	00000 - 59999

	35
	  Clr Tot. Key
	Modify
	3
	00000 - 59999

	36
	  Date –y/m/d  *
	Modify
	3
	99/12/31

	37
	  Time-h/m/s  *
	Modify
	3
	23/59/59

	38
	  Password L1
	Modify
	3
	0000 - 9999

	39
	  Password L2
	Modify
	3
	0000 - 9999

	40
	  Password L3
	Modify
	3
	0000 - 9999

	41
	  Current Zero
	Modify
	4
	0.0000 - 1.9999

	42
	  Current Max
	Modify
	4
	0.0000 - 3.9999

	43
	  Meter Factor
	Modify
	4
	0.0000 - 3.9999

	44
	  Convtr S/N
	Modify
	4
	0000000000-9999999999

	45
	  Sys Reset
	Password
	4
	


* Item No. 36 and 37 are optional and only effective for the converter with real clock and power failure recording function. 
6.5 Parameter Setting Explaination

6.5.1 Sensor Size

The converter supports sensor diameter ranging from 3 to 3000mm, which can be chosen by pressing UP or DOWN key.

6.5.2 Flow Range

Flow range refers to the upper range value (URV) of flow rate. The URV is relative to flow percentage and output signal. At the analog output the amount of the measured values in the range 0 up to URV is displayed linear to the current range 4 to 20mA, at the frequency output to the frequency range 0 to the end frequency. The low flow cutoff and flow limit alarm relates to flow range as well. The maximum measurable flow rate, however, is not limited to the flow range as long as the flow speed does not exceed 15m/s.

In this menu item, user can also choose unit of flow rate. For volume flow, L/s, L/min, L/h, m3/s, m3/min and m3/h are available; while for mass flow, kg/s、kg/m、kg/h、t/s、t/m、t/h can be selected from. It is up to the habits and application requirements to pickup a proper unit.

6.5.3 Auto Rng Chg

    The converter has a function called Auto-Range-Change that is usually used for control system with wide flow range variation. The primary flow range is the value given by menu item ‘Flow Range’. The second flow range (lower range) is obtained by selecting range ratio 1:2, 1:4 or 1:8 of primary one. 

    Fig 5 illustrates how the flow range is changed automatically. To safely change range and avoid vibration of display and output, a 5% to 10% hysteresis is added at the change point. 

[image: image6.wmf]Zero

Output

Flow

Hysteresis

5

% 

-

10

%

Down

-

going 

change point

Up

-

going Change 

Point

100

%

1

:

4 

Range ratio

Flow 

range 

1

Flow 

range 

2

4

mA

20

mA


Fig. 5 Illustration of Auto-Range-Change

6.5.4 Damping 

Long damping constant can improve the stability of display and output and is suitable to flow control application; while short damping constant has short response time and is suitable to the totalization of pulse flow. Damping time is selectable from 0.2s to 100s.

6.5.5 Flow Dir.
If the displayed direction sign is not agreed to the actual flow direction, change this item to the opposite option. 

6.5.6 Flow Zero
To conduct zero adjustment, the fluid in the sensor pipe must be held still. The flow zero is displayed by flow speed and the unit is m/s. The display of flow zero is shown below: 

      

On the LCD, the top line displays the measured zero point while the bottom line shows the adjustment value.  If the FS is not equal to 00.000m/s, adjust the sign and value on bottom line until FS back to nil. Remind again: to adjust the flow zero, the sensor pipe must be filled and the fluid must be kept still. The flow zero adjustment value is an important constant of the meter and should be printed on the calibration sheet and label. The value should include the sign and amount by unit of m/s.

6.5.7 L.F. Cutoff and Cutoff Enble
Low flow cutoff is set in percentage relative to flow range. If Cutoff is enabled and flow is lower than the set value, the display of flow rate, speed and percentage and signal outputs are forced to nil. If the item is disabled, no action is taken.

6.5.8 Rate-Of-Chng and Limit Time
‘Rate-of-change’ limit technique is used to eliminate application-related high electrical noise contained in the process flow signal. 

To check electrical noise, two parameters are defined: ‘Rate-of-change’ limit and ‘Control limit time’. If the sampled flow value exceeds the set rate-of-change limit value based on the averaged flow rate value up until the sampled time, the system will reject that sampled value and instead the averaged value including the rate-of-change limit value in place of the rejected sampled value will be output. However, if the limit-exceeding sampled value continues for the same flow direction for more than the preset control limit time, that data will be used as output signal. Fig 6 illustrates the effect of noise-suppressing by rate-of-change limit.

The value of rate-of-change limit can be set from 0 to 30% of flow range and limit time ranges from 0 to 20 seconds. If either of the two parameters is set to nil, the function is disabled.

The rate-of-change limit function is not suitable for short period measurement and flow meter calibration.

[image: image7.png]Unattenuated measuring variable

1= Control limit time tine

Flow with rate-of-change linit

O\

tine




Fig.6 Example for the effect of rate-of-change limit

6.5.9 Total Unit
The converter has three 10-digit counters and the maximum counts are 9999999999. The total flow unit can be L, m3, kg or t (metric ton) with a multiplying factor of 0.001, 0.01, 0.1, 1, 10, 100 or 1000.

6.5.10 Flow Density
    The converter is capable of measuring mass flow if fluid density is set. The density can be set from 0.0001 to 3.9999 and the mass unit is determined automatically by flow unit. The density should be set to 1.0000 (default value) if not used. Otherwise, measurement data will be forced to nil.

6.5.11 Current Type

Current output type is selectable from 4-20mA to 0-10mA. 

6.5.12 Pulse Output
Two types of pulse output are available to choose from: frequency output mode and pulse output mode. The meter outputs continuous square wave pulse under frequency mode, while pulse series under pulse mode. Frequency output is usually used for flow rate measurement and short period of time totalization. Pulse output can be connected to an external counter directly and is often used for long period of time totalization.

As mentioned hereinbefore, transistor open collector circuit is used for frequency and pulse output. Therefore, the external DC power supply and load are necessary. 
6.5.13 Pulse Factor
Pulse factor is defined as: volume or mass per pulse. It can be set to 0.001L/p, 0.01L/p, 0.1L/p, 1L/p, 2L/p, 5L/p, 10L/p, 100L/p, 1m3/p, 10 m3/p, 100 m3/p or 1000 m3/p. Pulse width is selectable from auto, 10ms, 20ms, 50ms, 100ms, 150ms, 200ms, 250ms, 300ms, 350ms and 400ms.

6.5.14 Freq Max
Frequency range corresponds to the upper range value of flow rate, or 100% of flow percentage in other word. Maximum frequency is selectable from 1 to 5999Hz.

6.5.15 Comm Address and Baudrate

    Substation address is needed when using RS485 communication. The address can be set from 01 to 99. Baud rate is the transmission speed between main and sub station. It is selectable from 600, 1200, 2400, 4800, 9600, 14400bps. Remind: the baud rate must be the same as that of the main computer.

6.5.16 EmpPipe Det.
This item is used to enable or disable the empty-pipe detector. If enabled, the meter will force the display value, analog output and digital output to nil when the sensor pipe is not full.

6.5.17 EmpPipe Alm.

This item is to set the electrode alarm trip value. Constant current source method is employed to measure the resistance between two electrodes. The variation of the resistance is checked by CPU and CPU recognizes if the pipe is empty or the electrodes are contaminated. The resistance is calculated as following: 


[image: image8.wmf]s

d

R

1

»


where, 

[image: image9.wmf]d

 = electrode radius




[image: image10.wmf]s

 = Fluid conductivity


The electrodes resistance is usually between 5 to 50kΩ。The variation of the resistance relates to the surface status of electrodes and variation of fluid characteristic. If the sensor is filled with fluid, abnormal resistance signal is detected and empty pipe alarm is output. 


The electrode alarm trip value is determined based on the first-time measured electrode resistance. After the installation of the flowmeter, measure the resistance between the electrodes when the sensor pipe is filled. Record the resistance value and take it as a basis. Usually, set the trip value as 3 times of the original resistance recorded. 

6.5.18 Hi ALM Enble
User can enable or disable the high limit alarm.

6.5.19 Hi Alm Limit
High alarm limit value is set in percentage of the upper range of flow rate. The parameter ranges from 0% to 199.9%. The meter outputs alarm signal when the flow percentage is higher than this value.

6.5.20 Lo Alm Enble
User can enable or disable the low limit alarm.

6.5.21 Lo Alm Limit
Low alarm limit value is set in percentage of the upper range of flow rate. The parameter ranges from 0% to 199.9%. The meter outputs alarm signal when the flow percentage is lower than this value.

6.5.22 Sensor S/N
Sensor serial number records the information of the sensor equipped with the converter and ensure them match up when installing. 

6.5.23 Sensor Fact.

The sensor factor is set according to the calibration sheet supplied by the manufacturer. Usually this factor has been set up by the manufacturer before shipping. It is an important value that determines the accuracy of measurement. Do not change it without calibration.

6.5.24 Field Mode
The converter offers three field exciting modes based on the exciting frequency. Mode 1 is the most-commonly used one and suitable for most cases. Mode 2 and 3 are low-frequency exciting modes and are better for large size meter to measure water. The calibration should be taken under the same exciting mode as that used for measurement.

6.5.25 RevMeas.Enbl: Reverse Measurement Enable
If RevMeas.Enbl is set to ON, the converter displays flow and outputs signals when flow direction is reversed. If OFF, the converter displays no flow and does not output signals when reversing.

6.5.26 Multiplying 

This item is a multiplying factor selectable from 0.0000 to 3.9999. When calculating the flow rate and total, this factor is taken into account. It is often used to measure the flow in the open channel. If not applied, set the value to 1.0000.

6.5.27 F. Total Set and R. Total Set

Presetting of forward and reverse total counter is designed to start counting from the existing reading when replacing a converter or flowmeter. It provides a continuous total flow read which is convenient for management.

6.5.28 Input Control
This menu item is set to select the function of contact input. There are three options to be chosen from: ‘input disabled’, ‘stop totalizer’ and ‘reset totalizer’. The converter disables the contact input if ‘input disabled’ is selected. The contact input is used to start/stop totalizer controlled by ON/OFF switch signal if ‘stop totalizer’ function is active. If ‘reset totalizer’ function is enabled, ON (close) contact signal will clear the three internal total flow counters.

6.5.29 Clr Totalizr 

Enter the ‘Totalizer Reset Password’ in this menu item and press ENTER to confirm. The converter clears the three internal counter and restart counting if password matched. 

6.5.30 Clr Tot. Key

The ‘Totalizer Reset Password’ is changeable in this menu item if Level-3 password is entered. Remind: keep the new password in a safe place.

6.5.31 Date –y/m/d and Time-h/m/s

   These items are used to change the internal real time clock if equipped. 

6.5.32 Password L1 ,Password L2 and Password L3

   To change the Level-1 to Level-3 passwords, use Level-4 or higher level password to enter and change these two items.

6.5.33 Current Zero and Current Max

    Adjust the current output zero point and upper range value. It is not suggested that user make any adjustment since it has been setup to the best condition by the manufacturer.

6.5.34 Meter Factor

This factor is used by the manufacturer to normalize the excitation current and amplifier signal of the converter. DO NOT change it.

6.5.35 Convtr S/N
This serial number records the manufacturing date and code of converter. DO NOT change it.

6.5.36 Sys Reset

This item is reserved for the manufacturer to re-initialize the converter. After system resetting, all settings are set to default values automatically.

7. INSTALLATION
The installation of flowmeter is shown as Fig. 7.




8. WIRING
8.1 GROUDING


Fig. 7 Grounding of Flowmeter and Pipe

8.2 Converter Terminals and Definition

   Terminal blocks and marks are shown in Fig. 8 and Fig. 9

[image: image11.emf] 

L

1

L

2

P

U

L

+

 

P

D

I

R

E

X

T

-

I

O

U

T

I

C

O

M

S

I

G

1

S

I

G

 

G

N

D

S

I

G

2

D

S

1

D

S

2

P

C

O

M

E

X

T

+

A

L

M

-

A

L

C

O

M

A

L

M

+

PE

PE

POWER

Ａ

Ｂ

I

N

+

I

N

-

A

2

+

V

D

B

2

R

T

S

G

N

D

SW1

ON

1 2


Fig. 8 Remote Type: Terminals and Marks

 The definition of terminals and their marks for remote type converter is given as below:

	DS1
	Shield drive 1

	SIG1
	Signal input 1

	SIG GND
	Signal Ground

	SIG2
	Signal input 2

	DS2
	Shield drive 2

	EXT+
	Coil excitation +

	EXT-
	Coil excitation -

	IOUT
	Current output +

	ICOM
	Current output -

	PUL+
	Frequency/pulse output +

	PCOM
	Frequency/pulse output -

	PDIR
	Flow direction indicator +

	ALM-
	Low alarm output +

	ALM+
	High alarm output +

	ALCOM
	Alarm output -

	A
	RS485 communication A

	B
	RS485 communication B

	IN+
	Input contact +

	IN-
	Input contact -

	L1(+)
	220V(24V +) input

	L2(-)
	220V(24V -) input


The dip switch SW1 is set to ON to supply +12V power to pulse output. If external power is used, turn the switch to OFF.

The definition of terminals and their marks for compact type converter is given as below:

	T -
	RS485-B

	T+
	RS485-A

	COM
	Alarm/flow direction/ pulse output -

	FDIR
	Flow direction indicator +

	AL
	Low alarm output +

	AH
	High alarm output +

	IN-
	Input contact -

	IN+
	Input contact +

	P+
	Frequency/pulse output +

	COM
	Current/pulse output -

	I+
	Current output +

	L1(+)
	220V(24V +) input

	L2(-)
	220V(24V -) input


8.3 Remote-type Wiring

8.3.1 Terminal Block in Sensor 

Fig. 11 Marks of Terminal Block


SIG1: Signal 1 (Connecting to white coax wire ofSTT3200 cable)

 
SIG2: Signal 2 (Connecting to black coax wire of STT3200 cable)


DS1: Signal 1 shield drive (Connecting to inner shield layer of white coax wire of STT3200 cable)

DS2: Signal 2 shield drive (Connecting to inner shield layer of black coax wire of STT3200 cable)

SIG GND: Signal ground (Connecting to Ex-shield of STT3200 cable)


EXT+: Coil 1 (Connecting to red cable)


EXT-: Coil 2 (Connecting to yellow cable)

8.3.2 Connection of STT3200 Cable

[image: image12.jpg]Black ey

Wite cable | |

VR

Sleeved grownd wire v}/

STT3200 Schematic Diag for Cable Preparation







Fig. 13 Schematic Diag for STT3200 Cable Preparation 

8.4 Output Signal Wiring

  The dip switch SW1 is set to ON to supply +12V power to pulse output. A 1KΩ resister is connected to the +12V power to provide a pull-up. If external power is used, turn the switch to OFF.

[image: image13.emf] 

IOUT

ICOM

Ammeter

－

＋


Fig. 14(a) Wiring of current output

[image: image14.emf] 

POUT+

PCOM

１

２ ３

４ ５ ６

Electromagnetic 

Counter

DC Power 

Supply

- +


Fig.14 (b) Example of electromagnetic counter connection


[image: image15.emf]１

２ ３

４ ５ ６

L1

L2

PCOM

IOUT+

IN+

PUL+


Fig. 14(c) Example of electrical counter connection

   
Fig. 14(d) Digital  Output Direct Connection  

 SHAPE \* MERGEFORMAT 


                                                 
 SHAPE \* MERGEFORMAT 



Fig.14 (f) Connection with relay (e.g. PLC)
Generally, the intermediate relay needs 12V or 24V power supply E. D is a surge-absorbing diode, which is usually embedded in the relay. If not, an external one is necessary.

9. Self-diagnostic and Troubleshooting

The converter is made by surface mount technology and is not repairable for user. Do not open the converter case.

 The self-diagnosis function of the converter is capable of displaying alarm information except power supply or hardware failures. A ‘!’ symbol is displayed on the right corner of LCD top-line and malfunction information can be read from the bottom-line by pressing DOWN key. User may check the flowmeter according to the alarm information. Some examples of alarms are given below: 








Coil Alm








Elctrd Alm








EpPipe Alm








Low Alarm








High Alarm

Troubleshooting information is given below:

9.1 No display

a) Check the connection of power supply;

b) Check fuse;

c) Check the voltage of power supply;

d) Check if the LCD contrast can be adjusted. Adjust it if possible;

e) Return to base, if a) to d) are OK.

9.2 Coil Alarm

a) Check if terminal EXT+ and EXT- are open;

b) Check if coil resistance is less than 150Ω;

c) Replace converter if a) and b) are OK.

9.3 Empty Pipe Alarm and Electrodes Alarm

a) Check if the sensor pipe is filled with fluid;

b) Check the connection of signal wiring;

c) Connect the terminal SIG1, SIG2 and SIG GND. If the alarm display disappears, it is confirmed the converter is normal. The alarm may be caused by the bubble in the fluid;

d) For electrodes alarm, measure the resistance between two electrodes with a multimeter. The read should be between 3 to 50kΩ. Otherwise, the electrodes are contaminated or covered.

9.4 High Alarm

Increase the flow range. 
9.5 Low Alarm
Reduce the flow range. 

9.6 Inaccurate Measurement

a) Check if the sensor pipe is filled with the fluid to be measured.

b) Check the wiring;

c) Check if the sensor factor and flow zero are the same as those on the calibration sheet.
10.  PACKAGING
The package includes: 

The electromagnetic flowmeter ordered;

Instruction Manual;

11. TRANSPORTATION AND STORAGE
To prevent the flowmeter from damage in the transportation, the package should be kept in unopened status before reaching installation site. The storage room should be satisfied with the following conditions:

a. Rain-proof, humidity-proof;

b. Strong Vibration and Shake Avoidance
c. Temperature between -20 to +60℃, relative humidity less than 80%

12.   OPERATION
Before operation, the following inspection should be done to check if:

a. There is any damage caused by transportation or installation;

b. The power used is same as the label on the flowmeter;

c. The wiring is correct.

After inspection, turn the valve on to fill the pipe up and make sure there is no leakage and the gas inside the pipe is eliminated. Switch on the power supply and the flowmeter is ready to use after 10 minutes warm-up.

If there is any problem, please refer to the Section 9 for troubleshooting. If still not working properly, contact the manufacturer immediately.

ELECTROMAGNETIC FLOWMETER

USER MANUAL

Jiang Su  Suke  Instrument  Co., Ltd
CONTENTS 

1、SUMMARY 
1

2. WORKING PRINCIPLES
1

2.1 Measuring Principles
1

  2.2 Converter Circuit Schematic
2
3. PRODUCT CLASSIFICATION
3

3.1 Product Components
3

3.2 Product Models
3
4. TECHNICAL SPECIFICATON
3

 5. APPEARANCE AND INSTALLATION
4

       6. CONVERTER OPERATION MENU AND PARAMETER SETUP
6
   6.1 Keypad
6
  6.2 Keypad Function
6
  6.3 Parameter Configuration and Operation Password
7
  6.4 Table of Parameter Setting Menu
7 

         6.5 Parameter Setting Explanation
8

       7. INSTALLATION
13
 8. WIRING
15
  8.1 Grounding
15
   8.2 Converter Terminals and Definition
16

 8.3 Remote-type Wiring
17 
   8.4 Output Signal Wiring
18
       9. SELF-DIAGNOSTICS AND TROUBLESHOOTING
20

 10. PACKAGING
21
 11. TRANSPORTATION AND STORAGE
21

 12. OPERATION
21




�





Fig3(a) Remote-type converter





Fig3(b) Compact-type converter











Remark: 


Nominal Pressure 4.0 MPa for DN10-150; Nominal Pressure 1.6 MPa for DN10-1000; Nominal Pressure 1.0 MPa for DN10-1600; Nominal Pressure 0.6 MPa for DN10-2000. Other special pressure level can be designed on demand.


The length L increases 6 mm if ground flange is equipped; The length L increases 6 mm if inlet protection flange is equipped; The length L increases 6 mm if liner protection flange is equipped.








+





282





.





74





m





3





/





h





P       





+





100





.





0  





%





∑





+





00123456





.





78 





m





3





Flow Velocity





Flow Percentage





Electrode Resistence 





Forward Total Flow





Reverse Total Flow





Net Total Flow





Alarm Display





ENTER





. 





Combine with  ALT to enter menu





.





DOWN





, 





  -





1





, 





PageDn





. 





Combine with ALT to shift left





ALT





, 





Combine with other keys





Flow Range Ratio





Flow Rate





Flow Unit





UP





, 





 + 1 











, 





PageUp





. 





Combine with ALT to shift right





Alarm Indicator





FS=○○.○○○m/s


±○○○○○








�





�





�





�





Fig. 9 Compact Type: Terminals and Marks











�





 Fig. 14(e) Connection with photo-coupler (e.g. PLC)











Circuit





Internal 





R





E





-





+





Device





User 





PCOM











PUL+





PUL+











PCOM





+





-





E





D





J





Internal Circuit 














18
1

_1234567891.unknown

_1234567892.unknown

_1234567893.vsd
１


２


３


４


５


６


L1


L2


PCOM


IOUT+


IN+


　


PUL+



_1234567890.unknown

